## Tabelle1
Bosmal Tender BOS-47-FZ-15
Position in RFQ | Text | Question | Answer | Question | Explanation
§2 A 2.4.2 | It is recommended that the sample of diluted exhaust gas is held at a temperature of 47 ± 5°C for ≥ 0.2 seconds immediately before meeting the filter face. | What is the technical reason for the measuring of the temperature ≥ 0.2 seconds before filter? Depends on the flow of the particulate sampling pump the particle is still outsite of the PM tunnel probe. Or is here a type error and the meaning is within 20cm according EPA CFR §1065.140 "Dilution for gaseous and PM constituents"? | This requirements refer to a heated cabinet and heated sampling probe for the particulate sampling system as supplementary scope §2 B 6. In basic scope it is not required. | Zwracamy się ponownie z prośbą o odpowiedź na pytanie dotyczące utrzymania próbki w temperaturze 47±5°C przez ≥0,2s bezpośrednio przed filtrami. Pytanie zaznaczone zostało w załączonym pliku na czerwono. Państwa odpowiedź nie wyjaśnia merytorycznie zadanego pytania, jest jedynie stwierdzeniem, że to wymaganie odnosi się do grzanej szafy oraz grzanej sondy próbkującej dla systemu cząstek stałych i nie jest wymagane w dostawie podstawowej. Prosimy o odpowiedź. | Zamiast zdania: "Rekomendowane jest utrzymywanie próbki w temperaturze 47±5°C przez ³0,2 s bezpośrednio przed filtrami." Powinno być zdanie: "Rekomendowane jest utrzymywanie próbki w temperaturze 47±5°C w odległości 20 cm od filtrów." Instead the sentence: „It is recommended that the sample of diluted exhaust gas is held at a temperature of 47 ± 5°C for ≥ 0.2 seconds immediately before meeting the filter face.” Should be the sentence: „It is recommended that the sample of diluted exhaust gas is held at a temperature of 47 ± 5°C within 20 cm upstream or downstream of the filter.”